## Slide 1
Normal Use, General Wear | Preventable Damage
Cosmetic items | Broken screw bosses or stand-offs
Non-skid pads worn or missing | Mechanical components damaged or missing
Velcro worn or missing | Housing damage
Loose hardware | Mechanical part fabrication or replacement
Screws missing | Connector housing damage
Coil stiffeners worn or broken | Connector contact damage(pins,etc.)
Insert stripped or missing | Cable trauma
Foam worn,broken-down | Foam torn
Let's Do Some Failure Analysis
· Compare the reported problem to main or contributing findings
Mechanical Condition

## Slide 2
Normal Use, Random Electronic Failures |  | Preventable Damage
Capacitor failure | Blown fuse | Connector housing damage
Diode failure | Intermittent system cable | Connector contact damage (pins,etc.)
Resistor failure | Intermittent flex circuits | Cable trauma
Inductor failure | Worn/Intermittent wiring | Broken solder joints
SCR failure | Internal wires | Broken interconnects
FET failure |  | Component failure due to trauma
RF switch failure |  | 
Preamp failure |  | 
Logic control failure |  | 
EEPROM failure |  | 
Flectronic tuning |  | 
Compare the reported problem to main or contributing findings
Electrical Analysis

## Slide 3
GE 1.5T Neurovascular Array
Part numbers:5117092-2,M3087JJ,M64NVA8-HD
>150 repairs
Most frequent failures
1. Intermittent solder connections and physical damage to electrical components
2. Cut,pull or damage to cable or strain relief
3. Random component failures
Normal Use/Wear
Preventable Damage
Electronic Worn wiring Intermittent
tuning flex traces
Intermittent wiring
Component Fail
Connector
trauma
Cable trauma
Component Fail

## Slide 4
Siemens 1.5T Body Flex Array
Part numbers:7579555,5369554
>150 repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Bent/damaged pins in connector
3. Intermittent solder connections and physical damage to electrical components
Intermittent flex traces
Intermittent Wiring
Connector Trauma
Component Fail due to
Cable Trauma
Components fail
Electronic tuning
Worn Wiring

## Slide 5
GE 1.5T 4Ch Small Flex Coil
Part number:5430692
>130 Repairs
Most frequent failures
1. Bent/damaged pins in connector
2. Intermittent/damaged wiring
3. Random component failures
Intermittent flex traces
Intermittent Wiring
Connector Trauma
Component Fail due to
Cable Trauma
Components fail
Electronic tuning
Worn Wiring
Preventable Damage:60%

## Slide 6
Toshiba 1.5T Neurovascular Atlas
Part numbers:MJAH-137A,2000205,2000206,2000207,2000208,5465781
>125 Repairs
Most frequent failures
1. Intermittent solder connections and
physical damage to electrical components
2. Damaged interconnects
3. Bent/damaged pins in connector
Normal Use/Wear
Preventable Damage
Component failure due to
Intermittent
flex traces
Intermittent wiring
Component failure
Electronic
tuning
Connector
trauma
Cable trauma Worn
wiring

## Slide 7
GE 1 .5T Express 16E Head Neck Array
Part numbers:5746586,5407991,5420021,M50002HL
>125 Repairs
Most frequent failures
1. Damaged interconnects
2. Intermittent solder connections and physical damage to electrical components
3. Random component failures
Preventable Damage 53%
Intermittent flex
traces
Component
failure due to
Component failure
Worn wiring
Electronic tuning
Intermittent wiring

## Slide 8
GE 1.5T Express 16E Head Neck Array
Part numbers:5746586,5407991,5420021,M50002HL
> 125 Repairs
Most frequent failures
1. Damaged interconnects
2. Intermittent solder connections and physical damage to electrical components
3. Random component failures
Preventable Damage
Electronic tuning
Intermittent flex
traces
Component
failure due to
Component failure
Intermittent wiring
Worn wiring

## Slide 9
Toshiba 1.5T Shoulder Atlas Coil
Part numbers:MJAJ-177A,5464955
>90 Repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Bent/damaged pins in connector
3. Intermittent solder connections and physical damage to electrical components
Normal Use/Wear
Preventable Damage
Intermittent flex traces
Intermittent flex traces
Intermittent Wiring
Connector Trauma
Component Fail due to
Cable Trauma
Components fail
Worn Wiring

## Slide 10
GE 1 .5T Express 9E Anterior Array
Part numbers:5407990,5433103,5746587,M50002HM
>80 Repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Random component failures
3. Bent/damaged pins in connector
Intermittent flex traces
Component Fail due to
Cable Trauma
Components fail
Worn Wiring

## Slide 11
Toshiba 1.5T Body Atlas SPEEDER Coil
Part number:MJAB-167A
>70 Repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Bent/damaged pins in connector
3. Intermittent solder connections and physical damage to electrical components
Preventable Damage
77%
Intermittent flex traces
Intermittent Wiring
Connector Trauma
Component Fail due to
Components fail

## Slide 12
GE 1 .5T Express 9E Anterior Array
Part numbers:5407990,5433103,5746587,M50002HM
>80 Repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Random component failures
3. Bent/damaged pins in connector
Preventable Damage
Intermittent flex traces
Component Fail due to
Cable Trauma
Components fail
Worn Wiring

## Slide 13
Normal Use/Wear
Preventable Damage 77%
Toshiba 1.5T Body Atlas SPEEDER Coil
Part number:MJAB-167A
>70 Repairs
Most frequent failures
1. Intermittent flex traces and wiring
2. Bent/damaged pins in connector
3. Intermittent solder connections and physical damage to electrical components
Intermittent flex traces
Intermittent Wiring
Connector Trauma
Component Fail due to
Components fail

## Slide 14
>60 Repairs
Most frequent failures
1. Bent/damaged pins and physical damage to connector
2. Intermittent solder connections and physical damage to electrical components
3. Intermittent wiring
Normal Use/Wear
Preventable Damage 82%
Siemens 1.5T Head Matrix Array
Part numbers:7577732,5319252
Connector trauma
Intermittent Wiring
Component Fail due to
Cable Trauma
Components fail
Electronic tuning
Worn Wiring

## Slide 15
Siemens 1.5T Extremity TR Matrix
Part number:474SI-64F
>60 Repairs
Most frequent failures
1. Intermittent solder connections and physical damage to electrical components
2. Random component failures
3. Intermittent wiring
Normal Use/Wear
Preventable Damage 81%
Intermittent
wiring
Component failure due to
Component failure
Connector
trauma
Electronic tuning
Cable trauma
wiring
Worn

## Slide 16
10 coil models
> 1800 repairs
>300 coil models
>1700 repairs
Comparing failures of top-10 to remaining volume
Component Intermittent failure due to wiring
Electronic Connector
tuning trauma
Cable trauma

## Slide 17
1.Prevention...Observe usage
· Transportation practices
· Patient placement practices
· Cable routing practices
· Cleaning practices
Physical
Manipulation
6 Strategies to Minimize Failures and Costs
Internal Inspection
Education
Partnership
Visual
Inspection
Early
Detection
Prevention

## Slide 18
Transportation Practices
· Transport the coil by gripping and supporting it from the sides
· NEVER pick up the coil from a single section (anterior section).
· When transporting between rooms,
ALWAYS use a cart
6 Strategies to Minimize Failures and Costs
Physical
Manipulation
Education
Partnership
Visual
Inspection
Internal
Inspection
Early
Detection
Prevention

## Slide 19
Transportation Practices
· Flex coils should ALWAYS be cradled from the bottom,gripped at the center housing or carried by the designated handle
- The coil should NEVER be carried by the flexible portion,as there are sensitive
components contained within that can be damaged by doing so
· DO NOT carry coils by the cable
- Component damage may occur
Physical
Manipulation
6 Strategies to Minimize Failures and Costs
Internal Inspection
Visual
Inspection
Education
Partnership
Early
Detection
Prevention

## Slide 20
Patient Placement Practices
· ALWAYS remove the anterior portion of a coil by pulling straight-up
- If the anterior is removed at an angle,the connections between the anterior and
posterior sections will be damaged
· With flex coils, DO NOT adjust the position of a patient using the coil,as it may cause both physical and electro-mechanical
damage
Internal Inspection
6 Strategies to Minimize Failures and Costs
Physical
Manipulation
Visual
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 21
Cable Routing Practices
· DO NOT loop or cross the cable when in use
- Looped cables may cause RF coupling and may degrade image quality
· DO NOT permit the cable to be bent at an extreme angle
- Route the cable using soft angles
- Damage can occur if the cable is bent to an extreme degree
· ALWAYS unplug the coil prior to lowering the table
- Failure to do so may damage the coil,cable or
Physical
Manipulation
6 Strategies to Minimize Failures and Costs
Internal Inspection
Education
Partnership
Visual
Inspection
Early
Detection
Prevention

## Slide 22
Cleaning Practices
· Clean the coil with a cloth dampened in a solution of:
- 10%bleach and 90%tap water,or
- 30%isopropyl alcohol and 70%tap water
- ALWAYS consult the OEM or OEM manual for approved chemicals
-ALWAYS detach the coil from the scanner before attempting to clean it
- DO NOT submerge any part of the coil
- DO NOT spray/pour cleaning agent directly on to coil,internal circuitry could be damaged
- Electric shock may occur if the coil is attached to the system during cleaning or when it is still wet
Internal Inspection
6 Strategies to Minimize Failures and Costs
Physical
Manipulation
Visual
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 23
2.External Visual Inspection
· OEMs recommend USERS visually inspect prior to EVERY use
· Realistically,managers should assign one MRI technologist to inspect EVERY coil EVERY week
- Alternate technologists every week
Any concerns would be relayed to a service engineer
Physical
Manipulation
Education
Partnership
Internal
Inspection
Visual
Inspection
Early
Detection
Prevention

## Slide 24
Internal Inspection
Bent pins in connectors
· Connectors should slide together with ease
6 Tips to Minimize Failures and Costs
Physical
Manipulation
Visual
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 25
Cracked connectors and housings and damaged foam indicate trauma and:
· Possible internal mechanical damage
· Broken solder joints
· Intermittent connections
· Broken electrical components
6 Tips to Minimize Failures and Costs
Internal Inspection
Physical
Manipulation
Visual
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 26
3.Early Detection
· Users need to alert service engineers as soon as a visual flaw is identified,NOT wait until a failure OCCurs
Cost difference may be hundreds versus thousands of dollars
Internal Inspection
6 Tips to Minimize Failures and Costs
Physical
Manipulation
Visual
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 27
4.Education/Partnership
· Ask for 15-minutes during a department meeting
· Educate technologists on
- Costs associated with repairing/replacing coils
- Best practices for transport,cable routing and patient manipulation
- How to visually inspect a coil
- When to alert you of a finding
Physical
Manipulation
6 Tips to Minimize Failures and Costs
Internal Inspection
Education
Partnership
Visual
Inspection
Early
Detection
Prevention

## Slide 28
· Rotate the coil in all directions listening
for internal noise,rattling,or creaking
6 Tips to Minimize Failures and Costs
5.Physical manipulation
Physical
Manipulation
Education
Partnership
Visual
Inspection
Internal
Inspection
Early
Detection
Prevention

## Slide 29
6.Internal Inspection
· NOT RECOMMENDED
- May induce additional problems
- Requires significant time
Does allow for
- Inspection for broken solder joints
- Perform continuity tests on wiring and cables
6 Strategies to Minimize Failures and Costs
Physical
Manipulation
Education
Partnership
Visual
Inspection
Internal
Inspection
Early
Detection
Prevention

## Slide 30
6 Tips to Minimize Failures and Costs
Internal Inspection
· Broken solder joints
Physical
Manipulation
Education
Partnership
Visual
Inspection
Internal
Inspection
Early
Detection
Prevention

## Slide 31
6 Tips to Minimize Failures and Costs
Internal Inspection
· Broken components
Physical
Manipulation
Visual
Inspection
Internal
Inspection
Education Partnership
Early
Detection
Prevention

## Slide 32
6 Tips to Minimize Failures and Costs
Internal Inspection
· Missing components
Physical
Manipulation
Education
Partnership
Visual
Inspection
Internal
Inspection
Early
Detection
Prevention

## Slide 33
Summary
· Based on our data
- >50%of coil failures are due to trauma
- Only 19% are due to actual random electronic failures
· Focus efforts on prevention,early detection,and education
· Base your troubleshooting on visual and functional cues
· Let those with expertise,experience,proper testing equipment, and proven procedures restore your coil's performance

## Slide 34
Thank you
Need material for Repair？？？ www.rf-coaxial.com